## Slide 1: Physics of Magnetic Resonance Chapter 12
Biomedical Engineering
Dr. Mohamed Bingabr
University of Central Oklahoma

## Slide 2: Outline
Introduction
Microscopic Magnetization
Macroscopic Magnetization
Precession and Larmor Frequency
Transverse and Longitudinal Magnetization
RF Excitation
Relaxation
The Bloch Equation
Spin Echoes
Basic Contrast Mechanisms

## Slide 3: Magnetic Resonance Imaging
  A projection of the three-dimensional volume of the body onto a two-dimensional imaging surface.
  MRI is the imaging of hydrogen density in tissues.
  Advantage:
  1- High image quality
  2- risk-free imaging
  Disadvantage: high cost
  Uses: Assess Neurological effects of stroke, trauma or disease. Orthopedic scans for injuries and degeneration involving knees, shoulders, feet and ankles.

## Slide 4: Magnetic Resonance Imaging
  Magnetic Field Strength
  Earth: 25 to 65 microTesla
  External Magnetic field B0 = 0.5 to 7 Tesla
  Gradient Magnetic Coil = 10-60 mT/meter
  1 Tesla = 10,000 Gauss
x
y
z

## Slide 5: Microscopic Magnetization
MR images of (a) the head showing the brain, spinal column, tongue, and vocal tract; (b) the knee; (c) the ankle; (d) the liver; and (e) the lumbar spine.

## Slide 6: Microscopic Magnetization
a) and (b) show two images of a wrist fracture. (c) and (d) show two images of a stroke. (e) and (f) show two MR images of a multiple sclerosis patient.

## Slide 7
1.5 T vs. 3.0 T (3D MR Angiography)
1.5 T
3.0 T

## Slide 8: Microscopic Magnetization
Nucleus with either an odd atomic number or an odd mass number has an angular momentum , and spin.
Microscopic magnetic field has a magnetic moment vector μ.
: gyromagnetic ratio with unit radian per second per tesla
Hz/Tesla

## Slide 9: Microscopic Magnetization
No net magnetization in the absence of external magnetic source.

## Slide 10: Nuclear Magnetization
Positive
Orientation
(Lower Energy)
Negative
Orientation
(Higher energy)
Magnetic
Field (B0)
54o
126o
Spin quantum number ½ system

## Slide 11: Macroscopic Magnetization
Net Magnetization
B0
M
r = (x, y, z)

## Slide 12: Macroscopic Magnetization
After a while M(r, t) will reach its equilibrium value M0
PD = proton density per unit volume
T: temperature from absolute value
h: Planck’s constant
k: Boltzmann’s constnt
Why can't we measure the Magnetic field M ?

## Slide 13: MR Imaging
The value of an MR image at a given tissue voxel is determined by:
1- Tissue Property
a) Relaxation parameters T1 and T2
b) Proton density
2- Scanner imaging protocol to manipulate vector M
a) Pulse sequence
b) Magnetic gradient

## Slide 14: Motion of Gyroscope
Equation that describes the motion of a gyroscope is
where L(t) is the gyroscope’s angular momentum, r the radius from the fixed point of rotation, m the mass, and g earth gravity.

## Slide 15: Precession and Larmor Frequency
M(t) is a magnetic moment and it experiences a torque at the presence of a time-varying magnetic field B(t).
If B(t) is a static magnetic field equal B0 at the z direction and the initial magnetization M(0) equal M0 and oriented at an angle  relative to the z axis then the solution is:

## Slide 16: Precession and Larmor Frequency
These equations describe a precession of M(t) around B0 with a frequency called Larmor frequency.

## Slide 17: Is Larmor Frequency Constant?
Three sources for B0 fluctuation:
Magnetic field inhomogeneities
  Shimming the main magnet.
Magnetic susceptibility (magnetic property that decreases or increases the magnetic field within the material)
  :diamagnetic susceptibility and spatially variable.
Chemical shift (Change to Larmor frequency due to chemical environment)
  shift
  :Shielding constant (ppm part per million)

## Slide 18: Longitudinal and Transverse Magnetization
Transverse Magnetization
Longitudinal Magnetization
Mxy(t) is the signal measured to form the MRI images

## Slide 19: Rotating Frame
A reference frame (rotating frame) is a frame that rotates at the Larmor frequency vo.
Viewing the transverse and longitudinal magnetization signals from the rotating frame.
Merry Go Rounds

## Slide 20: NMR Signals
A coil placed next to the area need to image will experience magnetic field radiated from the subject. This magnetic field will induce electric voltage in the coil proportional to the transverse magnetic field (Faraday’s Law).
Vs is the volume of the sample
Br is the magnetic field produced by the transmitter (coil) to control the magnetization vector M. Assume object is homogeneous and coil produce uniform field Br. Mz change slowly so dMz/dt is zero.

## Slide 21: NMR Signals
Remember
The frequency v0 of V(t) will determine the location of the voxel in the body from which the NMR is radiating, and the magnitude will determine the density of the H atoms in the voxel.
The x-y components of the magnetic field Br are

## Slide 22: Maximizing the Magnitude of NMR Signals
The goal is to maximize the magnitude of the NMR signal:
is called the tip angle or flip angle.
Increasing to /2 will maximize V but will increase the time to obtain the NMR signal.
Increasing Vs will reduce the resolution.
Example:
If we want to double the resolution in all three dimensions, by how much should we increase the value of B0 if it was initially 1.5 tesla?

## Slide 23: RF Excitation
RF excitation is established by a pulse of alternating current running through an antenna (coil) surrounding the sample.
The antenna will radiate circularly polarized magnetic field B1(t) with Larmor frequency using quadrature RF coils.
When the frequencies of B1(t) and M(t) are the same then M(t) will be pushed away from the external magnetic field B0 by tip angle α.
The value of α depends proportionally on the strength of the pulse and the time duration.
When the magnetization vector M is aligned with the strong external vector B0 it is very hard to detect M by the RF antenna. RF excitation is the tool to push M vector away from the B0 in order to detect M.

## Slide 24: RF Excitation
The circularly RF excitation pulse B1(t)
Rotational plane
Common RF pulses are the (pi over 2) and the (the inversion pulse). The final tip angle after an RF excitation of duration is
For rectangular pulse
is the envelop of B1(t) and is the initial phase

## Slide 25: RF Excitation
Example
We apply an RF pulse to a sample of protons. The sample is in equilibrium with the B0 field in the +Z direction. We need to tip the magnetization vector M into the x-y plane in 3 ms. What should the strength of RF excitation be?

## Slide 26: Relaxation after α Pulse Excitation
At the end of the α pulse, M will precess in response to the presence of the main magnetic field B0. The received signal Mxy(t) will slowly decay due to transverse and longitudinal relaxations mechanisms.
Mxy
Mz
Mxy
Mz
Mxy
Mz
M
M
M
Mxy
Mz
M0
Time Pass
Time Pass
Time Pass
Longitudinal relaxation
Transvers relaxation

## Slide 27: Transverse (Spin-Spin) Relaxation
Transverse Relaxation (spin-spin relaxation)
Perturbations in the magnetic field due to other spins that are nearby causes some protons to momentarily speed up or slow down, changing their phase. As a result Mxy exponentially decay to zero.

## Slide 28: Transverse Relaxation
The decayed received signal in the antenna is called free induction decay (FID). The decay time constant is called transverse relaxation time T2.
T2 depends on the types of tissues (causes contrast).
Local perturbations in the static field B0 makes the actual decay time for FID, , shorter so < T2.

## Slide 29: T2 of Some Normal Tissue Types
Tissue | T2 (ms)
gray matter | 100
white matter | 92
muscle | 47
fat | 85
kidney | 58
liver | 43
Even though the actual decay time for FID is , the T2 decay still can be measured by special RF pulsing sequence called spin echoes. Spin echoes pulse exploit the latent magnetization coherence that last T2.

## Slide 30: Longitudinal (spin-lattice) Relaxation
The longitudinal relaxation time (T1) is the time it takes for the longitudinal magnetization Mz(t) to recovers back to its equilibrium value M0. Mz(t) rises exponentially the rising time T1 depends on the tissue property.
: longitudinal magnetization immediately after the pulse.
Mz
M0
M
Mxy
Mz(0+)
M
α
M
Time

## Slide 31: T1 and T2 for Different Tissues
250 ms < T1< 2500 ms
25 ms < T2 < 250 ms
5 T2 < T1 < 10 T2

## Slide 32: Examples
A sample is in equilibrium if there have been no external excitation for at least 3 times the largest T1 in the sample.
Example:
Suppose a sample is in equilibrium, and a /2 pulse is applied. What happens to the longitudinal magnetization of the sample?
Example:
Suppose a sample is in equilibrium, and an  pulse is applied. What are the transverse and longitudinal magnetizations of the sample, expressed in the rotating and non rotating frames?

## Slide 33: The Bloch Equations
Bloch equations describe the behavior of the magnetic spin at the presence of the forced magnetic fields and the relaxation behavior.

## Slide 34: Solution of Bloch Equations
After  pulse, the RF field B1 is shut down and only B0 is nonzero. Therefore Bx(t) = By(t) = 0 and Bz(t) = B0 ,
Verify that the transverse and longitudinal relaxation solution satisfy Bloch equations.

## Slide 35: Magnetization M for /2 Pulse
Example:
The Bloch equations describe the behavior of M in the laboratory frame. Find the equations for the x, y, and z components of M after a /2 pulse (in the x direction) satisfy the equations.

## Slide 36: Spin Echoes to Measure T2
Pure transverse relaxation, characterized by the time constant T2, is a random phenomenon characterized by tissue properties. A /2 RF pulse followed by  pulse elicit the spin echoes transverse signal Mxy(t).
Two mechanisms that causes spin echo amplitude to decrease:
1- Longitudinal relaxation T1
2- Phase of the coherent echo is never perfectly aligned.
The problem is FID decay at the rate of , which is much smaller than T2
TR > 2000msec,
TE > 80msec

## Slide 37: Spin Echoes to Measure T2
TE is the echo time
T2
TR > 2000msec, TE > 80msec

## Slide 38: Spin Echoes to Measure T2

## Slide 39: Example
Suppose two 1H isochromats are in different locations in a 1.5 T magnet, and the fractional difference in field strength is 20 ppm (parts per million).
How long will it take before these isochromats are 180o out of phase?
What will be their phase difference at TE/2 if the echo time is 4 ms?

## Slide 40: T1-Weighted Contrast Image
The image intensity is proportional to the time relaxation of the longitudinal component of magnetization.

## Slide 41: Spin-Spin Pulse to Measure T1
B0
Tr
T1
TR < 1000msec, TE < 30msec

## Slide 42: Spin-Spin Pulse to Measure T1
When the sample at equilibrium and excited by  pulse then
After T1 which is larger than 3 times T2, Mxy will be negligible but Mz(t) will have value
If at time TR=T1 the tissue is excited with another  pulse then the transverse magnetization

## Slide 43: Basic Contrast Mechanisms
The transverse magnetization Mxy(t) produces the measurable MR signal.
Tissue contrast in MRI determined by
Tissue properties PD, T2, and T1.
Characteristic of the externally applied excitations:
  Tip angle 
  Echo time TE
  Pulse repetition interval TR

## Slide 44: Weighted Images
Three images of the same slice through the skull. Contrast between the tissue types are classified as (a) PD-weighted, (b) Τ2-weighted, and (c) Τ1-weighted.

## Slide 45: Brain Tissue Parameters
Water 3000 3000
Liver 45 420
Fat 85 240

## Slide 46: PD-Weighted Contrast Image
The image intensity is proportional to the number of hydrogen nuclei in the sample.
Need to image the sample in equilibrium before the signal has a chance to decay from T2 effects.
RF signal: Long TR = 3500 ms, short TE = 17 ms, and  = /2.

## Slide 47: T2-Weighted Contrast Image
The image intensity is proportional to the transverse relaxation times of different tissues.
RF signal: Long TR = 3500 ms, long TE = T2 of the tissues being imaged.
GM and WM has small contrast with respect to each other and large contrast with respect to CSF.
WM is slightly darker than GM because its NMR signal has decayed slightly faster.

## Slide 48: T1-Weighted Contrast Image
RF signal: Short TR = 600 ms, short TE = 17 ms and  = /2.
TR falls in between the T1 values for GM and WM but is much smaller than that of CSF.
GM and WM will have recovered approximately two-thirds of their longitudinal magnetization, whereas CSF will have recovered relatively little.
CSF signal smaller than GM and WM signals.
GM and WM are relatively bright, while the CSF is dark in the image.

## Slide 49: Brain Tissue Parameters
Weighted Image | TR | TE | 
PD | Long 3500 ms | Short 17 ms | /2
T2 | Long 3500 ms | long T2 | 
T1 | Short 600 ms | Short 17 ms | /2

## Slide 50: Inversion Recovery
Inversion recovery uses a 180o RF pulse to establish T1 contrast.
Let tnull be the time when Mz = 0
tnull = T1 ln 2
After 180o RF pulse
After /2 pulse at tnull Mz = 0 and
Tissues of different T1 will have Mxy values and will be imaged. Inversion recovery is used to suppress certain tissues

## Slide 51: Animation and More Detail
MRI is taught as a one semester course. For more detail and animation visit the following website:
http://www.cis.rit.edu/htbooks/mri/inside.htm
https://www.youtube.com/watch?v=Ok9ILIYzmaY
https://www.youtube.com/watch?v=zf5oX01bRgk
Magnetic Resonance Laboratory at Rochester Institute of Technology.

## Slide 52
Problem 12.6

## Slide 53
Problem 12.11

## Slide 54
Problem 12.14