## Slide 1: Signals and Systems Chapter 2
Biomedical Engineering
Dr. Mohamed Bingabr
University of Central Oklahoma

## Slide 2: Outline
Signals
Systems
The Fourier Transform
Properties of the Fourier Transform
Transfer Function
Circular Symmetry and the Hankel Transform

## Slide 3: Introduction
Signal Type
Continuous Signal: x-ray attenuation
Discrete Signal: times of arrival of photons in a radioactive decay process in PET
Mixed signal: CT scan signal g(l,θk)
System Type
Continuous-continuous system
  Continuous input  Continuous output
Continuous-discrete system
  Continuous input  Discrete output

## Slide 4: Signals
function
image
(x,y) : is a pixel location
f : is pixel intensity
2-D continuous signal is defined as f(x,y)

## Slide 5
Point Impulse
1-D point impulse (delta, Dirac, impulse function)
2-D point impulse
Point impulse is used in the characterization of image resolution and sampling

## Slide 6
Point Impulse Properties
1- Sifting property
We can interpret the product of a function with a point impulse as another point impulse whose volume is equal to the value of the function at the location of the point impulse.
2- Scaling property
2- Even function

## Slide 7
Line Impulse
This is a line whose unite normal is oriented at an angle θ relative to the x-axis and is at distance l from the origin in the direction of the unit normal.
The line impulse associated with line
Line also used to assist image resolution

## Slide 8
Comb and Sampling Functions
2-D comb function
Used in medical imaging production (sampling CT image 1024 x 1024), manipulation, and storage.
Sampling function

## Slide 9
1-D Rect and Sinc Functions
Rect function is used in medical imaging for sectioning.
Sinc function is used in medical imaging for reconstruction.

## Slide 10
2-D Rect and Sinc Functions

## Slide 11
Exponential and Sinusoidal Signals
x and y have distance units.
u0 and v0 are the fundamental frequencies and their units are the inverse of the units of x and y.

## Slide 12
Separable and Periodic Signals
A signal f(x, y) is separable if f(x, y)= f1(x) f2(y)
Separable signal model signal variations independently in the x and y direction.
Decomposing a signal to its components f1(x) and f2(y) might simplify signal processing.
Periodicity
A signal f(x, y) is periodic if f(x, y)= f(x+X, y) = f(x, y+Y)
X and Y are the signal periods in the x and y direction, respectively.

## Slide 13
Systems
A continuous system is defined as a transformer Ϩ of an input continuous signal f(x,y) to an output continuous signal g(x,y).
Linear Systems
g(x, y)= Ϩ [f(x, y)]

## Slide 14
Impulse Response
If we know the system response to an impulse
then with linearity we can know the system response to any input.
is the system impulse response function or known as point spread function (PSF).
System output g() for any input f().

## Slide 15
Impulse Response
System output g() for any input f().

## Slide 16
Shift Invariance System
A system is shift invariant if an arbitrary translation of the input results in an identical translation in the output.
Then with linearity we can know the system response to any input.
Let the input
then the output
Ϩ []= h()
System response to a shifted impulse
g()= Ϩ []

## Slide 17
Linear Shift-Invariance (LSI) System
Linear shift-invariant (LSI) System Response
Convolution Integral representation of system response
Example: Consider a continuous system with input-output equation g(x,y) = xyf(x,y).
Is the system linear and shift-invariant?

## Slide 18
Connection of LSI Systems
Cascade
Parallel

## Slide 19
Connection of LSI Systems
Example: Consider two LSI systems connected in cascade, with Gaussian PSFs of the form:
where σ1 and σ2 are two positive constants.
What is the PSF of the system?

## Slide 20
Separable Systems
A 2-D LSI system with PSF h(x, y) is a separable system if there are two 1-D systems with PSFs h1(x) and h2(y), such that h(x,y) = h1(x)h2(y)
This PSF is separable

## Slide 21
Separable Systems
In practice it is easier and faster to execute two consecutive 1-D operations than a single 2-D operation.
For every y
For every x

## Slide 22
Stable Systems
A system is a bounded-input bounded-output (BIBO) stable system if
For bounded input
for every (x, y)
The output is bounded
and

## Slide 23
1-D Fourier Transform (time)
Continuous 1-D Fourier Transform
Discrete 1-D Fourier Transform
x(n) = [125 145 148 140 110]
X(k) = [668 -29.2 - j38 7.7 - j12.96 7.7 - j12.96 -29.2 - j38]
|X(k)| = [668 47.9 15.1 15.1 47.9]
Phase = [0 -127.5 -59.3 59.3 127.5]
Ts = 0.25 sec
fs = 1/Ts = 4 Hz
fmax = fs/2 = 2 Hz
Sig length (T) = (N-1)*Ts=1 sec
fres = 1/T = 1 Hz

## Slide 24
1-D Fourier Transform
1-D inverse Fourier transform
1-D Fourier transform
Example:
What is the Fourier transform of the
u is the spatial frequency

## Slide 25
Fourier Transform
The 2-D Fourier transform of f(x, y)
u and v are the spatial frequencies
The 2-D inverse Fourier transform of F(u, v)

## Slide 26
Fourier Transform
Magnitude (magnitude spectrum) of FT
Angle (phase spectrum) of the FT
Example: What is the Fourier transform of the point impulse ?

## Slide 27
Fourier Transform Pairs

## Slide 28
Examples of Fourier Transform
Example: What is the Fourier transform of
Answer:
If the spatial frequency u0 and v0 are zero then f(x,y) =1 and the spectrum F(u,v) will be .
Slow signal variation in space produces a spectral content that is primarily concentrated at low frequencies.

## Slide 29
Examples of Fourier Transform
Three images of decreasing spatial variation (from left to right) and the associated magnitude spectra [depicted as log(1 + |F(u, υ)|)].

## Slide 30
Examples of Fourier Transform
>> img1 = imread('\\PHYSICSSERVER\MBingabr\BiomedicalImaging\mri.tif');
>> imshow(img1)
>> size(img1)
ans = 256 256
>> FFT_img1 = fftshift(fft2(img1));
>> Abs_FFT_img1 = abs(FFT_img1)
>> surf(Abs_FFT_img1(110:140,110:140))
>> Log_Abs_FFT_img1=log10(1+Abs_FFT_img1);
>> surf(Log_Abs_FFT_img1(110:140,110:140))

## Slide 31
Properties of the Fourier Transform
Linearity
Properties are used in theory and application to simplify calculation.
Translation
If F(u,v) is the FT of a signal f(x, y) then the FT of a translated signal
is

## Slide 32
Properties of the Fourier Transform
Conjugation and Conjugate Symmetry
If F(u,v) is the FT of a signal f(x, y) then

## Slide 33
Properties of the Fourier Transform
Scaling
If F(u,v) is the FT of a signal f(x, y) and if
Example
Detectors of many medical imaging systems can be modeled as rect functions of different sizes and locations. Compute the FT of the following

## Slide 34
Properties of the Fourier Transform
Rotation
If F(u,v) is the FT of a signal f(x, y) and if
If f(x, y) is rotated by an angle , then its FT is rotated by the same angle.

## Slide 35
Properties of the Fourier Transform
Convolution
The Fourier transform of the convolution
f(x, y) * g(x, y) is
Example:
Convolution property simplify the difficult task of calculating the convolution in the spatial domain to multiplication in the frequency domain.
Find Fourier transform of the convolution f(x, y) * g(x, y)
0 < V  U

## Slide 36
Properties of the Fourier Transform
Product
The Fourier transform of the product f(x, y) g(x, y) is the convolution of their Fourier transforms.
Separable Product
If f(x, y)=f1(x)f2(y) then
where

## Slide 37
Separability of the Fourier Transform
The Fourier transform F(u,v) of a 2-D signal f(x, y) can be calculated using two simpler 1-D Fourier transforms, as follows:
For every y.
1)
For every x.
2)

## Slide 39
Transfer Function
The system’s transfer function (frequency response) H(u, v) is the Fourier transform of the system’s PSF h(x,y).
The inverse Fourier transform of the transfer function H(u, v) is the point spread function h(x,y).
The output G(u, v) of a system in response to input F(u, v) is the product of the input with the transfer function H(u, v) .

## Slide 40
Transfer Function
Example: Consider an idealized system whose PSF is
h(x,y) = (x-x0, y-y0). What is the transfer function H(u, v) of the system, and what is the system output g(x, y) to an input signal f(x, y).

## Slide 41
Low Pass Filter
c1 > c2

## Slide 42
Circular Symmetry
Often, the performance of a medical imaging system does not depend on the orientation of the patient with respect to the system. The independence arises from the circular symmetry of the PSF.
A 2-D signal f(x, y) is circularly symmetric if
fθ(x, y) = f(x, y) for every θ.

## Slide 43
Property of Circular Symmetry
f(x, y) is even in both x and y
F(u, v) is even in both u and v
| F(u, v) | = F(u, v)
F(u, v) = 0
f(x, y) = f(r) where
F(u, v) = F(q) where
f(r) and F(q) are one dimensional signals representing two dimensional signals

## Slide 44
Hankel Transform
The relationship between f(r) and F(q) is determined by Hankel Transform.
where J0(r) is the zero-order Bessel function of the first kind.
The nth-order Bessel function
for n = 0, 1, 2, …

## Slide 45
Hankel Transform
The inverse Hankel transform.
unit disk
jink function

## Slide 46
sinc and jink functions
Example
In some medical imaging systems, only spatial frequencies smaller than q0 can be imaged. What is the function having uniform spatial frequencies within the desk of radius q0 and what is its inverse Fourier transform.