## Slide 1: Image Quality Chapter 3
Biomedical Engineering
Dr. Mohamed Bingabr
University of Central Oklahoma

## Slide 2: Image Quality Factors
Contrast
Resolution
Noise
Artifacts
Distortion
Accuracy

## Slide 3: Contrast
  Differences between image intensity of an object and surrounding objects or background.
  How to quantify contrast for
  image f(x, y)?
  Modulation

## Slide 4: Output Image Modulation

## Slide 5: Modulation Transfer Function (MTF)
  The MTF quantifies degradation of contrast as a function of spatial frequency.
  For most medical imaging
Isotropic system

## Slide 6: Modulation Transfer Function (MTF)
  Example
  What can we learn about the contrast behavior of an imaging system with this MTF?
input
output
u = 0.3
u = 0.5
u = 0.7

## Slide 7: Modulation Transfer Function (MTF)

## Slide 8: MTF for Nonisotropic System
  The MTF for nonisotropic system (PSF changes with orientation) has an orientation-dependence response.
  For most nonisotropic medical imaging

## Slide 9: Local Contrast
  Detecting a tumor in a liver requires local contrast.
  Example
  Consider an image showing an organ with intensity I0 and a tumor with intensity It > I0.
  What is the local contrast of the tumor? If we add a constant intensity Ic > 0 to the image, what is the local contrast? Is the local contrast improved?

## Slide 10: Resolution
  The ability of a medical imaging system to accurately depict two distinct events in space, time, or frequency as separate.
  Resolution could be spatial, temporal, or spectral resolution.
  High resolution is equivalent to low smearing

## Slide 11: Line Spread Function (LSF)
  Line impulse (line source)
  A vertical line impulse through the origin (θ = 0; l = 0)
The output g(x, y) of isotropic system for the input f(x, y) is
  Line spread function (LSF)

## Slide 12: Line Spread Function (LSF)
  Line spread function l(x) is related to the PSF h(x, y)
  Since the PSF h(x, y) is isotropic then l(x) is symmetric
  l(x) = l(-x)
  The 1-D Fourier transform L(u) of the LSF l(x) is
L(u) = H(u, 0)
l(x)

## Slide 13: Full Width at Half Maximum (FWHM)
  The FWHM of LSF (or PSF) is used to quantify resolution of medical imaging.
  The FWHM is the (full) width of the LSF (or the PSF) at one-half its maximum value. FWHM is measured in mm.
  The FWHM equals the minimum distance that two lines (or point) must be separated in space in order to appear as separate in the recording image.

## Slide 14: Resolution & Modulation Transfer Fun.
  For a sinusoidal input the spatial resolution is 1/u.
  The spatial frequency of the output depends on MTF cutoff frequency uc.
  Example:
  The MTF depicted in the Figure becomes zero at spatial frequencies larger than 0.8 mm-1. What is the resolution of the system?

## Slide 15: Resolution & Modulation Transfer Fun.
  Two systems with similar MTF curves but with different cutoff frequencies will have different resolutions, where MTF with higher cutoff frequency will have better resolution.
  It is complicated to use MTF to compare the frequency resolutions of two systems with different MTF curves.

## Slide 16: Resolution & Modulation Transfer Fun.
  Example: Sometimes, the PSF, LSF, or MTF can be described by a mathematical function by either fitting observed data or by making simplifying assumptions about the shape. Assume that the MTF of a medical imaging system is given by.
  What is the FWHM of this system?
  MTF can be directly obtained from the LSF.
  for every u

## Slide 17: Subsystem Cascade
  If resolution is quantified by FWHM, then the FWHM of the overall system (cascaded subsystems) is determined by
  The overall resolution of the system is determined by the poorest resolution of the subsystems (largest Ri) .
  If contrast and resolution are quantified using the MTF, then the MTF of the overall system will be given by

## Slide 18: Subsystem Cascade
  Resolution can depend on spatial and orientation, such ultrasound images.
  The MTF of the overall system will always be less than the MTF of each subsystem.

## Slide 19: Subsystem Cascade
  Example 3.5: Consider a 1-D medical imaging system with PSF h(x) composed of two subsystems with Gaussian PSFs of the form
  What is the FWHM of this system?
  Hint: from example 2.4

## Slide 20: Resolution Tool (bar phantom)
  Line pairs per millimeter (lp/mm)

## Slide 21: Noise
  Noise is any random fluctuation in an image; noise generally interferes with the ability to detect a signal in an image.
  Source and amount of noise depend on the imaging method used and the particular medical imaging system at hand.
  Example of source of noise: random arrival of photon in x-ray, random emission of gamma ray photon in nuclear imaging, thermal noise during amplifying radio frequency in MRI.

## Slide 22: Noise
  Random Variables
  Different repetitions of an experiment may produce different observed values. These values is the random variable.
  Probability Distribution Function (PDF)

## Slide 23: Continuous Random Variables
  Probability density function (pdf)
  Expected Value (mean)
  Variance

## Slide 24: Uniform Random Variable
  Probability density function ( pdf )
  The probability distribution function (PDF)

## Slide 25: Gaussian Random Variable
  Probability density function (pdf)
  The probability distribution function

## Slide 27: Discrete Random Variables
  Probability mass function (pmf)
  The probability distribution function

## Slide 28: Poisson Random Variables
  Poisson Random Variable
  Used in radiographic and nuclear medicine to statically characterize the distribution of photons count.

## Slide 29: Poisson Random Variables
  Example
  In x-ray imaging, the Poisson random variable is used to model the number of photon that arrive at a detector in time t, which is a random variable referred to as a Poisson process and given that notation N(t). The PMF of N(t) is given by
  Where λ is called the average rate of the x-ray photons.
  What is the probability that there is no photon detected in time t?

## Slide 30: Exponential Random Variables
  Example
  For the Poisson process of previous example, the time that the first photon arrives is a random variable, say T.
  What is the pdf of a random variable T?
  This is the pdf of exponential random variable T

## Slide 31: Independent Random Variables
  It is usual in imaging experiments to consider more than one random variable at a time.
  The sum S of the random variable N1, N2, …, Nm is a random variable with pdf
  Random variables are not necessary independent
  When random variables are independent

## Slide 32: Independent Random Variables
  Example:
  Consider the sum S of two independent Gaussian random variables N1 and N2, each having a mean of zero and variance of σ2.
  What are the mean, variance, and pdf of the resulting random variable?

## Slide 33: Signal-to-Noise Ration (SNR)
  The output of a medical imaging system g is a random variable that consists of two components f (deterministic signal) and g (random noise).

## Slide 34: Amplitude SNR
  Example:
  In projection radiography, the number of photons G counted per unit area by an x-ray image intensifier follows a Poisson distribution. In this case we may consider signal f to be the average photon count per unit area (i.e., the mean of G) and noise N to be the random variation of this count around the mean, whose amplitude is quantified by the standard deviation of G.
  What is the amplitude SNR of such a system?

## Slide 35: Power SNR
  Example:
  If f(x, y) is the input to a noisy medical imaging system with PSF h(x, y), then output at (x, y) maybe thought of as a random variable G(x, y) composed of signal
  h(x, y)*f(x, y) and noise N(x, y), with mean µN(x, y) and variance .
  What is the power SNR of such a system?
Answer depends on the nature of the noise:
1- White noise
2- wide-sense stationary noise

## Slide 36: Differential SNR
  ft and fb are the average image intensities within the target and background, respectively.
  A is the area of the target.
  C is the contrast.
  Noise: random fluctuation of image intensity from its mean over an area A of the background.
  Expressing SNR in decibels dB
  SNR (in dB) = 20 x log10 SNR (ratio of amplitude)
  SNR (in dB) = 10 x log10 SNR (ratio of power)

## Slide 37: Differential SNR
Example: consider the case of projection radiography. We may take fb to be the average photon count per unit area in the background region around a target, in which case fb = λb, where λb is the mean of the underlying Poisson distribution governing the number of background photons count per unit area. Notice that, in this case, .
What is the average number of background photons counted per unit area, if we want to achieve a desirable differential SNR?

## Slide 38: Sampling
Given a 2-D continuous signal f(x, y), rectangular sampling generate a 2-D discrete signal fd(m, n), such that
for m, n = 0, 1, …
Δx and Δy are the sampling periods in the x and y directions, respectively.
The inverse 1/Δx and 1/Δy are the sampling frequencies in the x and y directions, respectively.
What are the maximum possible values for Δx and Δy such that f(x, y) can be reconstructed from the 2-D discrete signal fd(m, n)?

## Slide 39: Sampling
Aliasing : When higher frequencies “take the alias of” lower frequencies due to under-sampling.

## Slide 40: Sampling
Original chest x-ray image and sampled images,
(b) without, and (c) with anti-aliasing

## Slide 41: Signal Mode for Sampling
Sampling is the multiplication of the continuous signal f(x, y) by the sampling function
Use Fourier series to represent the periodic impulses.
Use frequency shifting property.

## Slide 42: Nyquist Sampling Theorem
Sampling rate in y =
Sampling rate in x =

## Slide 43: Anti-Aliasing Filters
The image is passed through low-pass filter to eliminate high frequency components and then it can be sampled at lower sampling rate. The sampling rate equals or more than twice the cutoff frequency of the low-pass filter. This way aliasing will be eliminated but the low pass filtering introduces blurring in the image.
Example
Consider a medical imaging system with sampling period Δ in both the x and y directions.
What is the highest frequency allowed in the images so that the sampling is free of aliasing? If an anti-aliasing filter, whose PSF is modeled as a rect function, is used and we ignored all the side lobes of its transfer function, what are the widths of the rect function?

## Slide 44: Problem 3.22

## Slide 45: Other Effects
Artifacts
The creation of image features that do not represent valid anatomical or functional objects.
Examples of artifacts in CT: (a) motion artifact, (b) star artifact, (c) ring artifact, and (d) beam hardening artifact

## Slide 46: Other Effects
Distortion is geometrical in nature and refers to the inability of a medical imaging system to give an accurate impression of the shape, size, and/or position of objects of interest.

## Slide 47: Accuracy
Accuracy of medical image is judged by its ability in helping diagnosis, prognosis, treatment planning, and treatment monitoring. Here “accuracy” means both conforming to truth (free from error) and clinical utility. The two components of accuracy are quantitative accuracy and diagnostic accuracy.

## Slide 48: Quantitative Accuracy
Quantitative Accuracy refers to the accuracy, compared with the truth, of numerical values obtained from an image.
Source of Error
1- bias: systematic error
2- imprecision: random error

## Slide 49: Diagnostic Accuracy
Diagnostic Accuracy refers to the accuracy of interpretations and conclusions about the presence or absence of disease drawn from image patterns.
Diagnostic accuracy in clinical setting
Sensitivity (true-positive fraction): fraction of patients with disease who the test calls abnormal.
Specificity (true-negative fraction): fraction of patients without disease who the test calls normal.

## Slide 50: Sensitivity and Specificity
a and b, respectively, are the number of diseased and normal patients who the test calls abnormal.
c and d, respectively, are the number of diseased and normal patients who the test calls normal.
Sensitivity
Specificity
Diagnostic Accuracy (DA)
b
c

## Slide 51: Maximizing Diagnostic Accuracy
Because of overlap in the distribution of parameters values between normal and diseased patients, a threshold must be established to call a study abnormal such that both sensitivity and specificity are maximized.
Choice of Threshold
Relative cost of error
Prevalence: is a statistical concept referring to the number of cases of a disease that are present in a particular population at a given time
PR

## Slide 52: Diagnostic Accuracy
Two other parameters in evaluating diagnostic accuracy:
Positive predictive value (PPV): fraction of patients called abnormal who actually have the disease.
Negative predictive value (NPV); fraction of persons called normal who do not have the disease.
PPV
NPV
c
b

## Slide 53: Diagnostic Accuracy is not Enough
Example:
Consider a group of 100 patients, among which 10 are diseased and 90 are normal. We simply label all patients as normal. Construct the contingency table for this test and determine the sensitivity, specificity, and diagnostic accuracy of the test.

## Slide 54: Problem 3.21